Sensitivity of MKS Cirrus 3 residual gas analyzer for the Urania-specific gas mixtures

Purpose:

The purpose of this work is to quantitavely describe the precision of MKS Cirrus 3 RGA for the analysis of the following gas mixtures:

Table 1: Gas mixtures that are predicted to be analyed by RGA

Standard deviation of RGA readings would be used as a metric, that quantitavely describe precision of RGA readings. For example, if PPM of specific gas is said to be 1000±10 ppm, it means that 68% of readings of sample, that contains exactly 1000ppm of this specific gas, would fall into 990-1010ppm interval

Description of software workflow:

In order to convert raw RGA readings into concentration of gases in PPM, following algoritm is used:

1) Ionic currents and chamber pressure are inquired from RGA. Chamber pressure effectively do not change over single scan, but it is used to make ionic curent readings less dependent on pressure in gas lines. Absolute value of ionic current is divided over chamber pressure; obtained value has units of ampere/pascal metric. This value will be further reffered to as IC

2) As different gases have different ionization rates (for example, helium produces ~60 times lower ionic current than argon in the same concentration due to much higher ionization energy and much smaller atom size); IC is divided over calibration factor for specific gas. This factors are the following:

CO2	O2	N2	CH4	Ar	He
198993.795665839	100000	146258.079565515	154101.018002445	281868.233650486	5354.42339958041

CO2_28
15945.8734482825

Table 2: Calibration factors for establishing concentration of invididual compounds in the system. Note the second coefficient for CO2 (CO2_28), it is used to compute overlap between CO2 secondary peak and N2 primary peak at M/Z=28

Raw concentration, further reffered to as RAW, is not a PPM, and it cannot be normalized to produce PPM value directly, due to the fact that RGA sensitivity varies from day to day (by approximately 5%) for the reasons that I am not aware of. However, those changes of sensitivity do not affect precision of PPM measurement; as all reaings during single scan are affected uniformally

For N2 RAW readings, CO2 contribution to IC for M/Z=28 peak is subtracted first. CO2 contribution to M/Z=28 IC peak is computed as:

IC (M/Z=44) * CO2 (M/Z=28) IC->RAW coefficient / CO2 (M/Z=44) IC->RAW coefficient

Formula 2: Contribution of CO2 to M/Z=28 IC reading

3) RAW is converted to PPM by division over all RAW’s sum and multiplying by 1’000’000

Sources of uncertainty:

The main source of imprecision for determining of compound’s concentration is imprecision of IC measurement by the RGA. Linear fit, obtained for four series of ionic current scans, shows that standard deviation of measured value is a linear function of real value’s magnitude:

M/Z = 28	M/Z = 32	M/Z = 40	M/Z = 5

5929588.07018127	355321.76172342	1056299.87336386	-231.044328725393
5840622.89511461	351345.244048736	1054824.55581221	-13.2018942013623
5959824.37208008	346909.043640688	1019642.49632378	791.119202750482
6016397.37520933	358380.236092803	1051017.50802261	-789.942254306611
6023067.93128688	352700.317956873	1039264.37730252	-164.459325396764
5994609.10329992	349364.23627405	1022673.59614679	322.528637251919
6057426.00174497	362225.551337343	1078429.1126761	394.884395046637
6015842.70295918	367048.878460347	1082299.9144285	-611.079777659718
6101679.5486447	365738.779950731	1063897.77638949	-406.492931026926
5950395.22083808	365239.847201283	1097090.60163883	-216.28218751118
6047869.33180387	363623.986736641	1081689.42175664	330.990864401222
6267858.75553661	374095.096337554	1111469.83988205	-249.661166140465
5879025.68324264	352959.27028187	1049913.05037003	172.428490280583
5893940.52091272	350938.301209784	1053160.91918902	-6.63209793946094
6076374.06524124	360724.531011041	1052753.28205619	-165.640864562408
6088797.08798636	358417.213269194	1031345.83807394	152.514132008442
5982696.62507628	353907.461197597	1051540.71131547	165.637222534336
6115764.15779119	359800.972977349	1045293.21756655	-1124.82045075342
5964404.38191571	364053.925720222	1073673.86304442	918.337904156657
6040380.34473399	370868.210219158	1095515.91354043	350.208365527809
6177730.12818649	370322.210262092	1079337.25771023	376.473847697362
6103955.96411109	364545.913593516	1080613.91307155	263.99265078302
6272696.68650109	374352.436134673	1098576.52869867	-59.8997550518632
6313034.58817812	375500.616670457	1105258.52530232	-52.652994517747

6046415.89760735	361182.668429476	1065649.25390343	6.13773686021461
122018.02173486	8367.53573349343	25748.1297002289	461.743344808627

Table 3: IC’s obtained during sampling of 1% O2 1% N2 1% CH4 1% CO2 1% He 95% Ar calibration bottle. Last two lines display mean and standard deviation of data sets above

STDEV vs MEAN for ionic current measurement

Graph 1: Linear correlation between standard deviations and mean values of four data sets presented in Table 3 above

As linear fit shows strong correlation (R2 = 0.999); expected standard deviation of IC measurements would be computed using following formula, where SLOPE and INTERCEPT are obtained from linear fit equation at Graph 1:

ΔIC = SLOPE*IC + INTERCEPT

Formula 2: Calculation of expected IC measurement standard deviation

In order to proceed to IC standard deviation measurements; PPM values in Table 1 need to be converted to respective IC values. In order to do so; they are multiplied by expected RAW to PPM ratio (see bullet point 2 in Description of software workflow for justification of this step), and by RAW->IC coefficients in Table 2;

During multiple scans of 999’999 PPM Ar bottle, mean RAW value was found to be 520.86. Therefore, RAW->PPM coefficient is 0.00052086.

During this procedure, following table was produced:

	CO2	O2	N2	CH4	Ar	He
V142	497.212936627802	0	17.7092393100202	5.20859979706476	0.20834399188259	0.520859979706476
V140	377.727657283136	0	107.349241817505	31.4078567763005	1.25006395129554	3.1772458762095
V235	115.064219836917	0	285.085938712604	109.685298826488	2.77253767197757	6.09197832264694
V237	232.138438335521	0	172.545806337344	110.643160329168	0.559403618204755	0.614614776053642
V461	0	0	395.693680563152	108.743583983179	4.94764894723182	11.4750662129134
V458	0	0	0	520.859979706476	0	0
V336	0	0	498.755723887693	1.56570509899767	6.07478994331663	14.4637607764691
V459	0	0	506.110787661128	0	0.0192718192491396	14.7299202260991
V460	0	0	98.7029661543772	86.7398541404789	335.480183469164	0
V418	0	0	281.953486794649	0	238.906492911827	0
V446	0	0	0	488.182787159631	32.6771925468452	0
V447	0	0	506.065472842893	0	14.7945068635827	0
V412	0	0	0	0	520.859979706476	0
FS-V02	0	0	0	0	520.859979706476	0

Table 4: Expected RAW readings for gas mixtures encountered at Urania plant

Later, Table 4 is converted to table of predicted IC’s by multiplying the values in the table by coefficients in Table 2.

CO2 (M/Z = 44)	O2 (M/Z = 32)	N2 + CO2 (M/Z=28)	CH4 (M/Z = 15)	Ar (M/Z = 40)	He (M/Z = 4)
98942289.5137246	0	10518613.8963655	802650.531095008	58725.552983637	2788.90486324533
75165460.2507365	0	21723891.3719954	4839982.7025029	352353.317901822	17012.3196657965
22897065.8506766	0	43530921.3951823	16902616.2090642	781490.296329749	32619.0312805174
46194108.9643257	0	28937868.4321634	17050223.6417326	157678.109761065	3290.90773862949
0	0	57873397.8153769	16757496.9930423	1394585.06947892	61442.3630421579
0	0	0	80265053.1095008	0	0
0	0	72947054.3481222	241276.749647159	1712290.31112039	77445.0991474596
0	0	74022791.8507067	0	5432.11365098642	78870.229532578
0	0	14436106.2771592	13366699.8244314	94561206.7391944	0
0	0	41237975.5053861	0	67340151.1646891	0
0	0	0	75229464.4725701	9210662.54383608	0
0	0	74016164.1924159	0	4170101.51736806	0
0	0	0	0	146813882.459092	0
0	0	0	0	146813882.459092	0

Table 5: Predicted IC readings for M/Z ratios of interest for gas mixtures encountered at Urania plant

Now, for each of those values, standard deviation may be predicted using Formula 2. As concentration of specific compound is a function of six variables ( IC of every individual gas in the mixture ); it’s standard deviation is expressed as following:

CO rsub{2} PPM = {CO rsub{2} IC * CO rsub{2} IC->RAW } over { CO rsub{2} IC * CO rsub{2} IC->RAW + CH rsub{4} IC * CH rsub{4} IC->RAW + Ar IC * Ar IC->RAW + He IC * He IC->RAW + O rsub{2} IC * O rsub{2} IC->RAW + N rsub{2} IC * N rsub{2} IC->RAW}
{∂ CO rsub{2} PPM} over {∂ Ar IC} = { - CO rsub{2} IC * CO rsub{2} IC->RAW } over { ( CO rsub{2} IC * CO rsub{2} IC->RAW + CH rsub{4} IC * CH rsub{4} IC->RAW + Ar IC * Ar IC->RAW + He IC * He IC->RAW + O rsub{2} IC * O rsub{2} IC->RAW + N rsub{2} IC * N rsub{2} IC->RAW )^2 } * Ar IC->RAW

CO rsub{2} PPM = { 1 - CH rsub{4} IC * CH rsub{4} IC->RAW - Ar IC * Ar IC->RAW - He IC * He IC->RAW - O rsub{2} IC * O rsub{2} IC->RAW - N rsub{2} IC * N rsub{2} IC->RAW } over { CO rsub{2} IC * CO rsub{2} IC->RAW + CH rsub{4} IC * CH rsub{4} IC->RAW + Ar IC * Ar IC->RAW + He IC * He IC->RAW + O rsub{2} IC * O rsub{2} IC->RAW + N rsub{2} IC * N rsub{2} IC->RAW}

{∂ CO rsub{2} PPM} over {∂ CO rsub{2} IC} = { - 1 + CH rsub{4} IC * CH rsub{4} IC->RAW + Ar IC * Ar IC->RAW + He IC * He IC->RAW + O rsub{2} IC * O rsub{2} IC->RAW + N rsub{2} IC * N rsub{2} IC->RAW } over { ( CO rsub{2} IC * CO rsub{2} IC->RAW + CH rsub{4} IC * CH rsub{4} IC->RAW + Ar IC * Ar IC->RAW + He IC * He IC->RAW + O rsub{2} IC * O rsub{2} IC->RAW + N rsub{2} IC * N rsub{2} IC->RAW )^2 } * CO rsub{2} IC->RAW

ΔCO rsub{2} PPM = sqrt{ ( {∂CO rsub{2} PPM} over {∂CO rsub{2} IC}*ΔCO rsub{2} IC)^2 + ( {∂CO rsub{2} PPM} over {∂CH rsub{4} IC}*ΔCH rsub{4} IC)^2 + ( {∂CO rsub{2} PPM} over {∂N rsub{2} IC}*ΔN rsub{2} IC)^2 + ( {∂CO rsub{2} PPM} over {∂O rsub{2} IC}*ΔO rsub{2} IC)^2 + ( {∂CO rsub{2} PPM} over {∂Ar IC}*ΔAr IC)^2 + ( {∂CO rsub{2} PPM} over {∂He IC}*ΔHe IC)^2}